Biosafety Standards Comparison

LBNL, May 2008

This reference document was complied by Lawrence Berkeley National Laboratory (LBNL) as a resource to compare select biosafety requirements and guidelines by topic from the following biosafety standards:

- Biosafety in Microbiological and Biomedical Laboratories, Centers for Disease Control (CDC) and National Institutes of Health (NIH), fourth (1999) and fifth (2007) editions

- Guidelines for Research Involving Recombinant DNA Molecules (NIH Guidelines), National Institutes of Health (NIH), Effective January 2002, Federal Register, November 19, 2001 (66FR57970)

- Laboratory Biosafety Manual, World Health Organization (WHO), Geneva, third edition (2004)

- 29 CFR 1910.1030, Bloodborne Pathogens, OSHA (as of 2008)

Topics are generally organized according to standard and special Biosafety Level 1 (BL1) and Biosafety Level 2 (BL2) laboratory practices, safety equipment, personal protective equipment, and facilities as outlined in the BMBL and NIH Guidelines.

Consult the LBNL Biosafety Manual for biosafety requirements at LBNL. This document is only a resource.

Laboratory Biosafety Level 1

| Topic | BMBL 4th Edition 1999 | BMBL 5th Edition 2007 | NIH Guidelines 2002 | WHO Laboratory Biosafety Manual 3rd Ed |
| --- | --- | --- | --- | --- |
| A. BL1 Standard Microbiological Practices |  |  |
| Access Control | Access to the laboratory is limited or restricted at the discretion of the laboratory director when experiments or work with cultures and specimens are in progress. | The laboratory supervisor must enforce the institutional policies that control access to the laboratory. | Access to the laboratory is limited or restricted at the discretion of the Principal Investigator when experiments are in progress. | Only authorized persons should be allowed to enter the laboratory working areas. Children should not be authorized or allowed to enter laboratory working areas. Laboratory doors should be kept closed. |
| Decontamination | All cultures, stocks, and other regulated wastes are decontaminated before disposal by an approved decontamination method such as autoclaving. Work surfaces are decontaminated at least once a day and after any sp ill of viable material. | Decontaminate all cultures, stocks, and other potentially infectious materials before disposal using an effective method. | Work surfaces are decontaminated once a day and after any spill of viable material. All contaminated liquid or solid wastes are decontaminated before disposal. | Contaminated liquids must be decontaminated (chemically or physically) before discharge to the sanitary sewer. An effluent treatment system may be required, depending on the risk assessment for the agent(s) being handled. |
| Eating, drinking, smoking, handling of contact lenses, applying cosmetics. | Eating, drinking, smoking, handling contact lenses, applying cosmetics are not permitted in the work areas | Eating, drinking, smoking, handling contact lenses, applying cosmetics must not be permitted in the laboratory. | Eating, drinking, smoking, and applying cosmetics are not permitted in the work area. | Eating, drinking, smoking, applying cosmetics and handling contact lenses is prohibited in the laboratory working areas. |
| Food Storage | Storing food for human use is not permitted in the work areas. Food is stored outside the work area in cabinets or refrigerators designated and used for this purpose only. | Storing food for human consumption must not be permitted in laboratory areas. Food must be stored outside the laboratory area in cabinets or refrigerators designated and used for this purpose. | Food may be stored in cabinets or refrigerators designated and used for this purpose only. | Storing human foods or drinks anywhere in the laboratory working areas is prohibited. |
| Hand-washing | Persons wash their hands after they handle viable materials, after removing gloves, and before leaving the laboratory. | Hand washing protocols must be rigorously followed. Persons must wash their hands after working with potentially hazardous materials and before leaving the laboratory. | Persons wash their hands: (i) after they handle materials involving organisms containing recombinant DNA molecules and animals, and (ii) before exiting the laboratory. | Personnel must wash their hands after handling infectious materials and animals, and before they leave the laboratory working areas. |
| Hygiene |  |  | In the interest of good personal hygiene, facilities (e.g., hand washing sink, shower, changing room) and protective clothing (e.g., uniforms, lab coats) shall be provided that are appropriate for the risk of exposure to viable organisms containing recombinant DNA molecules. |  |
| Insect & Rodent Control | An insect and rodent control program is in effect. | An effective integrated pest management program is required | See Special Practices | None |
| Mouth Pipeting | Mouth pipetting is prohibited; mechanical pipetting devices are used. | Mouth pipetting is prohibited; mechanical pipetting devices must be used. | Mechanical pipetting devices are used; mouth pipetting is prohibited. | Pipetting by mouth must be strictly forbidden. |
| Sharps | Policies for the safe handling of sharps are instituted. | Policies for the safe handling of sharps, such as needles, scalpels, pipettes, and broken glassware must be developed and implemented. Whenever practical, laboratory supervisors should adopt improved engineering and work practice controls that reduce risk of sharps injuries. |  | The use of hypodermic needles and syringes should be limited. They must not be used as substitutes for pipetting devices or for any purpose other than parenteral injection or aspiration of fluids from laboratory animals. |
| Sharps |  | Precautions, including those listed below, must always be taken with sharp items. These include: a. Careful management of needles and other sharps are of primary importance. Needles must not be bent, sheared, broken, recapped, removed from disposable syringes, or otherwise manipulated by hand before disposal. b. Used disposable needles and syringes must be carefully placed in conveniently located puncture-resistant containers used for sharps disposal. c. Non disposable sharps must be placed in a hard walled container for transport to a processing area for decontamination, preferably by autoclaving. d. Broken glassware must not be handled directly. Instead, it must be removed using a brush and dustpan, tongs, or forceps. Plasticware should be substituted for glassware whenever possible. |  |  |
| Shoes | None | None | None | Open-toed footwear must not be worn in laboratories. |
| Splashes and Aerosols | All procedures are performed carefully to minimize the creation of splashes or aerosols. | Perform all procedures to minimize the creation of splashes and/or aerosols. | All procedures are performed carefully to minimize the creation of aerosols | All technical procedures should be performed in a way that minimizes the formation of aerosols and droplets. |
| Signs | A biohazard sign may be posted at the entrance to the laboratory whenever infectious agents are present. The sign may include the name of the agent(s) in use and the name an d phone number of the investigator. | A sign incorporating the universal biohazard symbol must be posted at the entrance to the laboratory when infectious agents are present. The sign may include the name of the agent(s) in use, and the name and phone number of the laboratory supervisor or other responsible personnel. Agent information should be posted in accordance with the institutional policy. | None | None |
| Spills | None | None | None- | All spills, accidents and overt or potential exposures to infectious materials must be reported to the laboratory supervisor. A written record of such accidents and incidents should be maintained. |
| Training | None | The laboratory supervisor must ensure that laboratory personnel receive appropriate training regarding their duties, the necessary precautions to prevent exposures, and exposure evaluation procedures. Personnel must receive annual updates or additional training when procedural or policy changes occur. | None | Continuous in-service training in safety measures is essential. |
| Work Surfaces | Work surfaces are decontaminated at least once a day and after any spill of viable material. | Decontaminate work surfaces after completion of work and after any spill or splash of potentially infectious material with appropriate disinfectant. | Work surfaces are decontaminated once a day and after any spill of viable material. | Work surfaces must be decontaminated after any spill of potentially dangerous material and at the end of the working day. |
| Transport of contamin-ated materials | Materials to be decontaminated outside of the immediate laboratory are packaged in accordance with applicable local, state, and federal regulations before removal from the facility. Materials to be decontaminated outside of the immediate laboratory are to be placed in a durable, leak-proof container and closed for transport from the laboratory. | Materials to be removed from the facility for decontamination must be packed in accordance with applicable local, state, and federal regulations. Depending on where the decontamination will be performed, the following methods should be used prior to transport: Materials to be decontaminated outside of the immediate laboratory must be placed in a durable, leak proof container and secured for transport. | See Special Practices below | To avoid accidental leakage or spillage, secondary containers, such as boxes, should be used, fitted with racks so that the specimen containers remain upright. The secondary containers may be of metal or plastic, should be auto-clavable or resistant to the action of chemical disinfectants, and the seal should preferably have a gasket. They should be regularly decontaminated. Packing and transportation must follow applicable national and/or international regulations. |
| B. BL1 Special Practices |  |  |  |
|  | None | None | Contaminated materials that are to be decontaminated at a site away from the laboratory are placed in a durable leak-proof container which is closed before being removed from the lab. An insect and rodent control program is in effect. | None |
| C. BL1 Safety Equipment (Primary Barriers and PPE) |  |  |
| Biosafety Cabinets | Special containment devices or equipment such as biological safety cabinets are generally not required for manipulations of agents assigned to BSL 1. | Special containment devices or equipment, such as BSCs, are not generally required. | Special containment equipment is generally not required for manipulations of agents | None |
| Lab Coats, Gowns, or Uniforms. | It is recommended that lab coats, gowns, or uniforms be worn to prevent contamination or soiling of street clothes. | Protective laboratory coats, gowns, or uniforms are recommended to prevent contamination of personal clothing. | None | Laboratory coveralls, gowns or uniforms must be worn at all times for work in the laboratory. |
| Eye wash | None | None | None | Safety systems should cover fire, electrical emergencies, emergency shower and eyewash facilities. |
| Eye Wear | Persons who wear contact lenses in laboratories should also wear goggles or a face shield. | Wear protective eyewear when conducting procedures that have the potential to create splashes of microorganisms or other hazardous materials. Persons who wear contact lenses in laboratories should also wear eye protection. | None | Safety glasses, face shields (visors) or other protective devices must be worn when it is necessary to protect the eyes and face from splashes, impacting objects and sources of artificial ultraviolet radiation |
| Protective Laboratory Coats, Gowns, or uniforms. | Protective eyewear should be worn for conduct of procedures in which splashes of micro-organisms or other hazardous materials is anticipated. | Protective laboratory coats, gowns, or uniforms are recommended to prevent contamination of personal clothing. | None | .None |
| Gloves | Gloves should be worn if the skin on the hands is broken or if a rash is present. Alternatives to powdered latex gloves should be available. | Gloves must be worn to protect hands from exposure to hazardous materials. Glove selection should be based on an appropriate risk assessment. Alternatives to latex gloves should be available. Wash hands prior to leaving the laboratory. In addition, BSL-1 workers should: Change gloves when contaminated, integrity has been compromised, or when otherwise necessary. b. Remove gloves and wash hands when work with hazardous materials has been completed and before leaving the laboratory. Do not wash or reuse disposable gloves. Dispose of used gloves with other contaminated laboratory waste. | None | Appropriate gloves must be worn for all procedures that may involve direct or accidental contact with blood, body fluids and other potentially infectious materials or infected animals. After use, gloves should be removed aseptically and hands must then be washed. |
| Backflow | None | None | None | A dependable supply of good quality water is essential. There should be no cross-connections between sources of laboratory and drinking-water supplies. An anti-backflow device should be fitted to protect the public water system. |
| Bench Tops | Bench tops are impervious to water and are resistant to moderate heat and the organic solvents, acids, alkalis, and chemicals used to decontaminate the work surface and equipment. | Bench tops must be impervious to water and resistant to heat, organic solvents, acids, alkalis, and other chemicals. | Bench tops are impervious to water and resistant to acids, alkalis, organic solvents, and moderate heat. | Bench tops should be impervious to water and resistant to disinfectants, acids, alkalis, organic solvents and moderate heat. |
| Carpets and Rugs | Carpets and rugs in laboratories are not appropriate. | Carpets and rugs in laboratories are not appropriate. | None | None |
| Cleanliness | The laboratory is designed so that it can be easily cleaned. | The laboratory should be designed so that it can be easily cleaned. Spaces between benches, cabinets, and equipment should be accessible for cleaning | The laboratory is designed so that it can be easily cleaned. | The laboratory should be kept neat, clean and free of materials that are not pertinent to the work. Open spaces between and under benches, cabinets and equipment should be accessible for cleaning. |
| Chairs | None | Chairs used in laboratory work must be covered with a non-porous material that can be easily cleaned and decontaminated with appropriate disinfectant. | None | Constructed of materials that are impermeable to liquids, resistant to corrosion and meet structural requirements |
| Doors | Laboratories should have doors for access control. | Laboratories should have doors for access control. | None | Doors should have vision panels, appropriate fire ratings, and preferably be self-closing. |
| Eating | None | None | None | Facilities for eating and drinking and for rest should be provided outside the laboratory working areas. |
| Electrical | None | None | None | There should be a reliable and adequate electricity supply and emergency lighting to permit safe exit. A stand-by generator is desirable for the support of essential equipment, such as incubators, biological safety cabinets, freezers, etc., and for the ventilation of animal cages. |
| Gas | None | None | None | There should be a reliable and adequate supply of gas. Good maintenance of the installation is mandatory. |
| Laboratory Furniture | Laboratory furniture is capable of supporting anticipated loading and uses. | Laboratory furniture must be capable of supporting anticipated loads and uses. | Laboratory furniture is sturdy. | Laboratory furniture should be sturdy. |
| Lighting | None | None | None | Illumination should be adequate for all activities. Undesirable reflections and glare should be avoided. |
| Security | None | None | None | Laboratories and animal houses are occasionally the targets of vandals. Physical and fire security must be considered. Strong doors, screened windows and restricted issue of keys are compulsory. Other measures should be considered and applied, as appropriate, to augment security |
| Sink | Each laboratory contains a sink for hand-washing. | Laboratories must have a sink for hand washing. | Each laboratory contains a sink for hand washing. | Hand-washing basins, with running water if possible, should be provided in each laboratory room, preferably near the exit door. |
| Spaces | Spaces between benches, cabinets, and equipment are accessible for cleaning. | None | Spaces between benches, cabinets, and equipment are accessible for cleaning. | Storage space must be adequate to hold supplies for immediate use and thus prevent clutter on bench tops and in aisles. Additional long-term storage space, conveniently located outside the laboratory working areas, should also be provided. |
| Storage | None | None | None | Space and facilities should be provided for the safe handling and storage of solvents, radioactive materials, and compressed and liquefied gases |
| Ventilation | None | None | None | In the planning of new facilities, consideration should be given to the provision of mechanical ventilation systems that provide an inward flow of air without re-circulation. |
| Windows | If the laboratory has windows that open to the exterior, they are fitted with fly screens. | Laboratories windows that open to the exterior should be fitted with screens. | If the laboratory has windows that open, they are fitted with fly screens. | When windows can be opened, they should be fitted with arthropod-proof screens. |
| Walls | None | None | None | Walls, ceilings and floors should be smooth, easy to clean, impermeable to liquids and resistant to the chemicals and disinfectants normally used in the laboratory. Floors should be slip-resistant. |

Laboratory Biosafety Level 2

| Topic | CDC BMBL 4th Edition 1999 | CDC BMBL 5th Edition 2007 | NIH Guidelines 2002 | OSHA Bloodborne Pathogens Standard 29 CFR 19101.1030 | WHO Laboratory Biosafety Manual 3rd Edition 2004 |
| --- | --- | --- | --- | --- | --- |
| A. BL2 Standard Microbiological Practices |  |  |  |
| Access Control | Access to the laboratory is limited or restricted at the discretion of the laboratory director when experiments are in progress. The laboratory director has the final responsibility for assessing each circumstance and determining who may enter or work in the laboratory or animal room. | The laboratory supervisor must enforce the institutional policies that control access to the laboratory. | Access to the laboratory is limited or restricted by the Principal Investigator when work with organisms containing recombinant DNA molecules is in progress. The Principal Investigator limits access to the laboratory. The Principal Investigator has the final responsibility for assessing each circumstance and determining who may enter or work in the laboratory. | HIV and HBV Research Laboratories and Production Facilities: Laboratory doors shall be kept closed when work involving HIV or HBV is in progress. Access to the work area shall be limited to authorized persons. | Only authorized persons should be allowed to enter the laboratory working areas. Laboratory doors should be kept closed. Children should not be authorized or allowed to enter laboratory working areas. Access to animal houses should be specially authorized. |
| Decontamination | All cultures, stocks, and other regulated wastes are decontaminated before disposal by an approved decontamination method such as autoclaving. Work surfaces are decontaminated on completion of work or at the end of the day and after any spill or splash of viable material with disinfectants that are effective against the agents of concern. | Decontaminate all cultures, stocks, and other potentially infectious materials before disposal using an effective method. Depending on where the decontamination will be performed, the following methods should be used prior to transport: Decontaminate work surfaces after completion of work and after any spill or splash of potentially infectious material with appropriate disinfectant. Laboratory equipment should be routinely decontaminated, as well as, after spills, splashes, or other potential contamination. Decontaminate work surfaces after completion of work and after any spill or splash of potentially infectious material with appropriate disinfectant. | Work surfaces are decontaminated at least once a day and after any spill of viable material. All contaminated liquid or solid wastes are decontaminated before disposal. | HIV and HBV Research Laboratories and Production Facilities: Before disposal all waste from work areas and from animal rooms shall either be incinerated or decontaminated by a method such as autoclaving known to effectively destroy bloodborne pathogens. | Contaminated liquids must be decontaminated (chemically or physically) before discharge to the sanitary sewer. An effluent treatment system may be required, depending on the risk assessment for the agent(s) being handled. Written documents that are expected to be removed from the laboratory need to be protected from contamination while in the laboratory. |
| Disposal | None | Eye and face protection must be disposed of with other contaminated laboratory waste or decontaminated before reuse. Persons who wear contact lenses in laboratories should also wear eye protection. | None | Disposal of all regulated waste shall be in accordance with applicable regulations of the United States, States and Territories, and political subdivisions of States and Territories. Regulated waste shall be placed in containers which are: Closable; Constructed to contain all contents and prevent leakage of fluids during handling, storage, transport or shipping; Labeled or color-coded in accordance with paragraph (g)(1)(i) this standard; and closed prior to removal to prevent spillage or protrusion of contents during handling, storage, transport, or shipping. All personal protective equipment shall be removed prior to leaving the work area. When personal protective equipment is removed it shall be placed in an appropriately designated area or container for storage, washing, decontamination or disposal. Regulated waste that has been decontaminated need not be labeled or color-coded. | All contaminated materials, specimens and cultures must be decontaminated before disposal or cleaning for reuse. Waste is anything that is to be discarded. In laboratories,decontamination of wastes and their ultimate disposal are closely interrelated. In terms of daily use, few if any contaminated materials will require actual removal from the laboratory or destruction. Most glassware, instruments and laboratory clothing will be reused or recycled. The overriding principle is that all infectious materials should be decontaminated, autoclaved or incinerated within the lab |
| Eating, drinking, smoking, handling of contact lenses, applying cosmetics | Eating, drinking, smoking, handling contact lenses, and applying cosmetics are not permitted in the work areas. | Eating, drinking, smoking, handling contact lenses, applying cosmetics must not be permitted in laboratory areas. | Eating, drinking, smoking, and applying cosmetics are not permitted in the work area. | Eating, drinking, smoking, applying cosmetics or lip balm, and handling contact lenses are prohibited in work areas where there is a reasonable likelihood of occupational exposure. | Eating, drinking, smoking, applying cosmetics and handling contact lenses is prohibited in the laboratory working areas. |
| Food Storage | Food is stored outside the work area in cabinets or refrigerators designated for this purpose only. | Storing food for human consumption must not be permitted in laboratory areas. Food must be stored outside the laboratory area in cabinets or refrigerators designated and used for this purpose. | Food may be stored in cabinets or refrigerators designated and used for this purpose only. | Food and drink shall not be kept in refrigerators, freezers, shelves, cabinets or on countertops or benchtops where blood or other potentially infectious materials are present. | Storing human foods or drinks anywhere in the laboratory working areas is prohibited. |
| Handwashing | Persons wash their hands after they handle viable materials, after removing gloves, and before leaving the laboratory. | Persons must wash their hands after working with potentially hazardous materials and before leaving the laboratory. | Persons wash their hands: (i) after handling materials involving organisms containing recombinant DNA molecules and animals, and (ii) when exiting the laboratory. | When provision of handwashing facilities is not feasible, the employer shall provide either an appropriate antiseptic hand cleanser in conjunction with clean cloth/paper towels or antiseptic towelettes. When antiseptic hand cleansers or towelettes are used, hands shall be washed with soap and running water as soon as feasible. Employers shall ensure that employees wash their hands immediately or as soon as feasible after removal of gloves or other personal protective equipment. | Personnel must wash their hands after handling infectious materials and animals, and before they leave the laboratory working areas. After use, gloves should be removed aseptically and hands must then be washed. |
| Insect & Rodent Control | An insect and rodent control program is in effect. | An effective integrated pest management program is required. | An insect and rodent control program is in effect. |  | There should be an arthropod and rodent control programme. |
| Mouth Pipetting | Mouth pipetting is prohibited; mechanical pipetting devices are used. | Mouth pipetting is prohibited; mechanical pipetting devices must be used. | Mechanical pipetting devices are used; mouth pipetting is prohibited. | Mouth pipetting/ suctioning of blood or other potentially infectious materials is prohibited. | Pipetting by mouth must be strictly forbidden. Pipetting aids – to avoid mouth pipetting. Many different designs are available. Materials must not be placed in the mouth. Labels must not be licked. |
| Sharps | Policies for the safe handling of sharps are instituted. Plasticware should be substituted for glassware whenever possible. | Policies for the safe handling of sharps, such as needles, scalpels, pipettes, and broken glassware must be developed and implemented. Whenever practical, laboratory supervisors should adopt improved engineering and work practice controls that reduce risk of sharps injuries. | Hypodermic needles and syringes are used only for parenteral injection and aspiration of fluids from laboratory animals and diaphragm bottles. | Contaminated needles and other contaminated sharps shall not be bent, recapped, or removed except as noted in paragraphs (d)(2)(vii)(A) and (d)(2)(vii)(B) below. Shearing or breaking of contaminated needles is prohibited. | The use of hypodermic needles and syringes should be limited. They must not be used as substitutes for pipetting devices or for any purpose other than parenteral injection or aspiration of fluids from laboratory animals. |
| Sharps | Syringes which re-sheathe the needle, needleless systems, and other safety devices are used when appropriate. | Careful management of needles and other sharps are of primary importance. Needles must not be bent, sheared, broken, recapped, removed from disposable syringes, or otherwise manipulated by hand before disposal. | Only needle-locking syringes or disposable syringe needle units (i.e., needle is integral to the syringe) are used for the injection or aspiration of fluids containing organisms that contain recombinant DNA molecules. Extreme caution should be used when handling needles and syringes to avoid auto-inoculation and the generation of aerosols during use and disposal. Needles should not be bent, sheared, replaced in the needle sheath or guard, or removed from the syringe following use. The needle and syringe should be promptly placed in a puncture-resistant container and decontaminated, preferably autoclaved, before discard or reused. | Contaminated needles and other contaminated sharps shall not be bent, recapped or removed unless the employer can demonstrate that no alternative is feasible or that such action is required by a specific medical or dental procedure. | After use, hypodermic needles should not be recapped, clipped or removed from disposable syringes. The complete assembly should be placed in a sharps disposal container. Disposable syringes, used alone or with needles, should be placed in sharps disposal containers and incinerated, with prior autoclaving if required. |
| Sharps |  | Used disposable needles and syringes must be carefully placed in conveniently located puncture-resistant containers used for sharps disposal. |  | Immediately or as soon as possible after use, contaminated reusable sharps shall be placed in appropriate containers until properly reprocessed. These containers shall be: Puncture resistant; Labeled or color-coded in accordance with this standard; Leakproof on the sides and bottom. Reusable sharps that are contaminated with blood or other potentially infectious materials shall not be stored or processed in a manner that requires employees to reach by hand into the containers where these sharps have been placed. | Sharps disposal containers must be puncture-proof/-resistant and must not be filled to capacity. When they are three-quarters full they should be placed in “infectious waste” containers and incinerated, with prior autoclaving if laboratory practice requires it. Sharps disposal containers must not be discarded in landfills. |
| Sharps |  | Non-disposable sharps must be placed in a hard walled container for transport to a processing area for decontamination, preferably by autoclaving. |  | Such bending, recapping or needle removal must be accomplished through the use of a mechanical device or a one-handed technique. |  |
| Sharps |  | Broken glassware must not be handled directly. Instead, it must be removed using a brush and dustpan, tongs, or forceps. Plasticware should be substituted for glassware whenever possible. | . | HIV and HBV Research Laboratories and Production Facilities: Hypodermic needles and syringes shall be used only for parenteral injection and aspiration of fluids from laboratory animals and diaphragm bottles. Only needle-locking syringes or disposable syringe-needle units (i.e., the needle is integral to the syringe) shall be used for the injection or aspiration of other potentially infectious materials. Extreme caution shall be used when handling needles and syringes. A needle shall not be bent, sheared, replaced in the sheath or guard, or removed from the syringe following use. The needle and syringe shall be promptly placed in a puncture-resistant container and autoclaved or decontaminated before reuse or disposal. |  |
| Shoes | None | None | None | Surgical caps or hoods and/or shoe covers or boots shall be worn in instances when gross contamination can reasonably be anticipated (e.g., autopsies, orthopaedic surgery). | Open-toed footwear must not be worn in laboratories. |
| Signs | A biohazard sign must be posted on the entrance to the laboratory when etiologic agents are in use. Appropriate information to be posted includes the agent(s) in use, the biosafety level, the required immunizations, the investigator’s name and telephone number, any personal protective equipment that must be worn in the laboratory, and any procedures required for exiting the laboratory. | A sign incorporating the universal biohazard symbol must be posted at the entrance to the laboratory when infectious agents are present. Posted information must include: the laboratory’s biosafety level, the supervisor’s name (or other responsible personnel), telephone number, and required procedures for entering and exiting the laboratory. Agent information should be posted in accordance with the institutional policy. | When the organisms containing recombinant DNA molecules in use in the laboratory require special provisions for entry (e.g., vaccination), a hazard warning sign incorporating the universal biosafety symbol is posted on the access door to the laboratory work area. The hazard warning sign identifies the agent, lists the name and telephone number of the Principal Investigator or other responsible person(s), and indicates the special requirement(s) for entering the laboratory. | HIV and HBV Research Laboratories and Production Facilities: When other potentially infectious materials or infected animals are present in the work area or containment module, a hazard warning sign incorporating the universal biohazard symbol shall be posted on all access doors. The hazard warning sign shall comply with paragraph (g)(1)(ii) of this standard. These signs shall be fluorescent orange-red or predominantly so, with lettering and symbols in a contrasting color. The employer shall post signs at the entrance to work areas which shall bear the following legend: (Name of the Infectious Agent) (Special requirements for entering the area) (Name, telephone number of the laboratory director or other responsible person.) | The international biohazard warning symbol and sign (Figure 1) must be displayed on the doors of the rooms where microorganisms of Risk Group 2 or higher risk groups are handled. |
| Space/Storage | None | None | Experiments of lesser biohazard potential can be conducted concurrently in carefully demarcated areas of the same laboratory. | None | Facilities for storing outer garments and personal items should be provided outside the laboratory working areas. Space and facilities should be provided for the safe handling and storage of solvents, radioactive materials, and compressed and liquefied gases. |
| Splashes and Aerosols | All procedures are performed carefully to minimize the creation of splashes or aerosols. | Perform all procedures to minimize the creation of splashes and/or aerosols. | All procedures are performed carefully to minimize the creation of aerosols | All procedures involving blood or other potentially infectious materials shall be performed in such a manner as to minimize splashing, spraying, spattering, and generation of droplets of these substances. | All technical procedures should be performed in a way that minimizes the formation of aerosols and droplets. |
| Training | The laboratory director ensures that laboratory and support personnel receive appropriate training on the potential hazards associated with the work involved, the necessary precautions to prevent exposures, and the exposure evaluation procedures. Personnel receive annual updates or additional training as necessary for procedural or policy changes. | The laboratory supervisor must ensure that laboratory personnel receive appropriate training regarding their duties, the necessary precautions to prevent exposures, and exposure evaluation procedures. Personnel must receive annual updates or additional training when procedural or policy changes occur The laboratory supervisor must ensure that laboratory personnel receive appropriate training regarding their duties, the necessary precautions to prevent exposures, and exposure evaluation procedures. Personnel must receive annual updates or additional training when procedural or policy changes occur. Personal health status may impact an individual’s susceptibility to infection, ability to receive immunizations or prophylactic interventions. Therefore, all laboratory personnel and particularly women of child-bearing age should be provided with information regarding immune competence and conditions that may predispose them to infection. Individuals having these conditions should be encouraged to self-identify to the institution’s healthcare provider for appropriate counseling and guidance. |  | Employers shall ensure that all employees with occupational exposure participate in a training program which must be provided at no cost to the employee and during working hours. Training shall be provided as follows: At the time of initial assignment to tasks where occupational exposure may take place; At least annually thereafter. Employers shall provide additional training when changes such as modification of tasks or procedures or institution of new tasks or procedures affect the employee's occupational exposure. The additional training may be limited to addressing the new exposures created. Material appropriate in content and vocabulary to educational level, literacy, and language of employees shall be used. | Staff training should always include information on safe methods for highly hazardous procedures that are commonly encountered by all laboratory personnel and which involve: 1. Inhalation risks (i.e. aerosol production) when using loops, streaking agar plates, pipetting, making smears, opening cultures, taking blood/serum samples, centrifuging, etc. 2. Ingestion risks when handling specimens, smears and cultures 3. Risks of percutaneous exposures when using syringes and needles 4. Bites and scratches when handling animals 5. Handling of blood and other potentially hazardous pathological materials 6. Decontamination and disposal of infectious material. |
| Training |  | The laboratory supervisor must ensure that laboratory personnel demonstrate proficiency in standard and special microbiological practices before working with BSL-2 agents. |  | HIV and HBV Research Laboratories and Production Facilities: Additional training requirements for employees in HIV and HBV research laboratories and HIV and HBV production facilities are specified in paragraph (g)(2)(ix). | The laboratory supervisor (reporting to the laboratory director) should ensure that regular training in laboratory safety is provided. |
| Transportation | Materials to be decontaminated outside of the immediate laboratory are placed in a durable, leak-proof container and closed for transport from the laboratory. Materials to be decontaminated off-site from the facility are packaged in accordance with applicable local, state, and federal regulations, before removal from the facility. | Potentially infectious materials must be placed in a durable, leak proof container during collection, handling, processing, storage, or transport within a facility. Potentially infectious materials must be placed in a durable, leak proof container during collection, handling, processing, storage, or transport within a facility. Materials to be decontaminated outside of the immediate laboratory are placed in a durable, leak-proof container and secured for transport from the laboratory. |  | Specimens of blood or other potentially infectious materials shall be placed in a container which prevents leakage during collection, handling, processing, storage, transport, or shipping. HIV and HBV Research Laboratories and Production Facilities: Contaminated materials that are to be decontaminated at a site away from the work area shall be placed in a durable, leakproof, labeled or color-coded container that is closed before being removed from the work area. | Packing and transportation must follow applicable national and/or international regulations. |
| Transportation | Cultures, tissues, specimens of body fluids, or potentially infectious wastes are placed in a container with a cover that prevents leakage during collection, handling, processing, storage, transport, or shipping. |  |  | The container for storage, transport, or shipping shall be labeled or color-coded according to paragraph (g)(1)(i) and closed prior to being stored, transported, or shipped. When a facility utilizes Universal Precautions in the handling of all specimens, the labeling/color-coding of specimens is not necessary provided containers are recognizable as containing specimens. This exemption only applies while such specimens/containers remain within the facility. Labeling or color-coding in accordance with paragraph (g)(1)(i) is required when such specimens/ containers leave the facility. |  |
| Transportation |  |  |  | If outside contamination of the primary container occurs, the primary container shall be placed within a second container which prevents leakage during handling, processing, storage, transport, or shipping and is labeled or color-coded according to the requirements of this standard. |  |
| Transportation |  |  |  | If the specimen could puncture the primary container, the primary container shall be placed within a secondary container which is puncture-resistant in addition to the above characteristics. |  |
| B. BL2 Special Practices |  |  |  |
| Access Control (Immunization) | The laboratory director establishes policies and procedures whereby only persons who have been advised of the potential hazards and meet specific entry requirements (e .g., immunization) may enter the laboratory. | The Principal Investigator establishes policies and procedures whereby only persons who have been advised of the potential hazard and meet any specific entry requirements (e.g., immunization) may enter the laboratory or animal rooms. | The Principal Investigator establishes policies and procedures whereby only persons who have been advised of the potential hazard and meet any specific entry requirements (e.g., immunization) may enter the lab or animal rooms | Written policies and procedures shall be established whereby only persons who have been advised of the potential biohazard, who meet any specific entry requirements, and who comply with all entry and exit procedures shall be allowed to enter the work areas and animal rooms. |  |
| Access Warning | In general, persons who are at increased risk of acquiring infection, or for whom infection may have serious consequences, are not allowed in the laboratory or animal room s. For example, persons who are immunocomprom ised or immunosuppressed m ay be at increased risk of acquiring infections. | All persons entering the laboratory must be advised of the potential hazards and meet specific entry/exit requirements. | .None | Universal precautions shall be observed to prevent contact with blood or other potentially infectious materials. Under circumstances in which differentiation between body fluid types is difficult or impossible, all body fluids shall be considered potentially infectious materials. The employer shall ensure that this information is conveyed to all affected employees, the servicing representative, and/or the manufacturer, as appropriate, prior to handling, servicing, or shipping so that appropriate precautions will be taken.. | Personnel should be advised of special hazards, and required to read the safety or operations manual and follow standard practices and procedures. The lab supervisor should make sure that all personnel understand these. A copy of the safety or operations manual should be available in the laboratory. |
| Animals | Animals not involved in the work being performed are not permitted in the lab. | Animals and plants not associated with the work being performed must not be permitted in the laboratory. | Animals not involved in the work being performed are not permitted in the laboratory. |  | No animals should be admitted other than those involved in the work of the laboratory. |
| Aerosol Generation | Aerosol generation processes may include centrifuging, grinding, blending, vigorous shaking or mixing, sonic disruption, opening containers of infectious materials whose internal pressures may be different from ambient pressures, inoculating animals intranasally, and harvesting infected tissues from animals or embryonate eggs. | All procedures involving the manipulation of infectious materials that may generate an aerosol should be conducted within a BSC or other physical containment devices. Procedures with a potential for creating infectious aerosols or splashes are conducted in BSC. These may include pipetting, centrifuging, grinding, blending, shaking, mixing, sonicating, opening containers of infectious materials, inoculating animals intranasally, and harvesting infected tissues from animals or eggs. High concentrations or large volumes of infectious agents are used. | All procedures are performed carefully to minimize the creation of aerosols. Procedures with a high potential for creating aerosols are conducted (see Appendix G-III-O, Footnotes and References of Appendix G). These may include centrifuging, grinding, blending, vigorous shaking or mixing, sonic disruption, opening containers of materials whose internal pressures may be different from ambient pressures, intranasal inoculation of animals, and harvesting infected tissues from animals or eggs. | HIV and HBV Research Laboratories and Production Facilities: All activities involving other potentially infectious materials shall be conducted in biological safety cabinets or other physical-containment devices within the containment module. No work with these other potentially infectious materials shall be conducted on the open bench. | These may include centrifugation, grinding, blending, vigorous shaking or mixing, sonic disruption, opening of containers of infectious materials whose internal pressure may be different from the ambient pressure, intranasal inoculation of animals, and harvesting of infectious tissues from animals and eggs. Plastic disposable transfer loops. Alternatively, electric transfer loop incinerators may be used inside the biological safety cabinet to reduce aerosol production. |
| Biosafety Manual | Biosafety procedures are incorporated into standard operating procedures or in a biosafety manual adopted or prepared specifically for the laboratory by the laboratory director. Personnel are advised of special hazards and are required to read and follow instructions on practices and procedures. | A laboratory-specific biosafety manual must be prepared and adopted as policy. The biosafety manual must be available and accessible. | A biosafety manual is prepared or adopted. Personnel are advised of special hazards and are required to read and follow instructions on practices and procedures. | Each employer having an employee(s) with occupational exposure as defined by paragraph (b) of this section shall establish a written Exposure Control Plan designed to eliminate or minimize employee exposure. HIV and HBV Research Laboratories and Production Facilities: A biosafety manual shall be prepared or adopted and periodically reviewed and updated at least annually or more often if necessary. Personnel shall be advised of potential hazards, shall be required to read instructions on practices and procedures, and shall be required to follow them. | It is the responsibility of the laboratory director (the person who has immediate responsibility for the laboratory) to ensure the development and adoption of a biosafety management plan and a safety or operations manual. Each laboratory should adopt a safety or operations manual that identifies known and potential hazards, and specifies practices and procedures to eliminate or minimize such hazards |
| Decontamination | Laboratory equipment and work surfaces should be decontaminated with an effective disinfectant on a routine basis, after work with infectious materials is finished, and especially after overt spills, splashes, or other contamination by infectious materials. Contaminated equipment must be decontaminated according to any local, state, or federal regulations before it is sent for repair or maintenance or pack aged for transport in accordance with applicable local, state, or federal regulations, before removal from the facility. | Equipment must be decontaminated before repair, maintenance, or removal from the laboratory. Laboratory equipment should be routinely decontaminated, as well as, after spills, splashes, or oth*er potential contamination. Equipment must be decontamin- ated before repair, maintenance, or removal from the laboratory | All wastes from laboratories and animal rooms are appropriately decontaminated before disposal. | Cleaning, Laundering, and Disposal. The employer shall clean, launder, and dispose of personal protective equipment required by paragraphs (d) and (e) of this standard, at no cost to the employee. |  |
| Exposure Incidents | None | Incidents that may result in exposure to infectious materials must be immediately evaluated and treated according to procedures described in the laboratory biosafety safety manual. All such incidents must be reported to the laboratory supervisor. Medical evaluation, surveillance, and treatment should be provided and appropriate records maintained. | None | Following a report of an exposure incident, the employer shall make immediately available to the exposed employee a confidential medical evaluation and follow-up, HIV and HBV Research Laboratories and Production Facilities: A spill or accident that results in an exposure incident shall be immediately reported to the laboratory director or other responsible person. |  |
| Gloves | None | None | Special care is taken to avoid skin contamination with organisms containing recombinant DNA molecules; gloves should be worn when handling experimental animals and when skin contact with the agent is unavoidable. | Gloves shall be worn when it can be reasonably anticipated that the employee may have hand contact with blood, other potentially infectious materials, mucous membranes, and non-intact skin; when performing vascular access procedures except as specified in paragraph (d)(3)(ix)(D); and when handling or touching contaminated items or surfaces. Disposable (single use) gloves such as surgical or examination gloves, shall be replaced as soon as practical when contaminated or as soon as feasible if they are torn, punctured, or when their ability to function as a barrier is compromised. Disposable (single use) gloves shall not be washed or decontaminated for re-use. |  |
| Immunization and Medical Surveillance | Laboratory personnel receive appropriate immunizations or tests for the agents handled or potentially present in the laboratory (e.g., hepatitis B vaccine or TB skin testing). Laboratory personnel receive appropriate immunizations or tests for the agents ha ndled or potentially present in the laboratory (e.g., hepatitis B vaccine or TB skin tes ting). | Laboratory personnel must be provided medical surveillance and offered appropriate immunizations for agents handled or potentially present in the laboratory. Personal health status may impact an individual’s susceptibility to infection, ability to receive immunizations or prophylactic interventions. Therefore, all laboratory personnel and particularly women of child-bearing age should be provided with information regarding immune competence and conditions that may predispose them to infection. Individuals having these conditions should be encouraged to self-identify to the institution’s healthcare provider for appropriate counseling and guidance. | None | The employer shall make available the hepatitis B vaccine and vaccination series to all employees who have occupational exposure, and post-exposure evaluation and follow-up to all employees who have had an exposure incident. | Appropriate medical evaluation, surveillance and treatment should be provided for all personnel in case of need, and adequate medical records should be maintained. |
| Laundry | None | None | None | Contaminated laundry shall be handled as little as possible with a minimum of agitation. Contaminated laundry shall be bagged or containerized at the location where it was used and shall not be sorted or rinsed in the location of use. Contaminated laundry shall be placed and transported in bags or containers labeled or color-coded in accordance with paragraph (g)(1)(i) of this standard. When a facility utilizes Universal Precautions in the handling of all soiled laundry, alternative labeling or color-coding is sufficient if it permits all employees to recognize the containers as requiring compliance with Universal Precautions. Whenever contaminated laundry is wet and presents a reasonable likelihood of soak-through of or leakage from the bag or container, the laundry shall be placed and transported in bags or containers which prevent soak-through and/or leakage of fluids to the exterior. The employer shall ensure that employees who have contact with contaminated laundry wear protective gloves and other appropriate personal protective equipment. contaminated laundry off-site to a second facility which does not utilize Universal Precautions in the handling of all laundry, the facility generating the contaminated laundry must place such laundry in bags or containers which are labeled or color-coded in accordance with paragraph (g)(1)(i). |  |
| Record-Keeping | None | None | None | Training records shall include the following information: The dates of the training sessions; The contents or a summary of the training sessions; The names and qualifications of persons conducting the training; and The names and job titles of all persons attending the training sessions. Training records shall be maintained for 3 years from the date on which the training occurred | none |
| Serum Collection and Storage | When appropriate, considering the agent(s) handled, baseline serum samples for laboratory and other at-risk personnel are collected and stored. Additional serum specimens may be collected periodically, depending on the agents handled or the function o f the facility. | Each institution must establish policies and procedures describing the collection and storage of serum samples from at-risk personnel. | When appropriate, considering the agent(s) handled, baseline serum samples for laboratory and other at-risk personnel are collected and stored. Additional serum specimens may be collected periodically depending on the agents handled or the function of the facility. |  |  |
| Sharps | Needles and syringes or other sharp instruments should be restricted in the laboratory for use only when there is no alternative, such as parenteral injection, phlebotomy, or aspiration of fluids from lab animals and diaphragm bottles. Plasticware should be substituted for glassware whenever possible. |  |  |  |  |
| Sharps | Only needle-locking syringes or disposable syringe needle units (i.e., needle is integral to the syringe) are used for injection or aspiration of infectious materials. Used disposa ble need les mu st not be bent, sheared, broken, recapped, removed from disposable syringes, or otherwise manipulated by hand before disposal; rather, they must be carefully place in conveniently located puncture-resistant containers used for sharps disposal. Non-disposable sharps must be placed in a hard-walled container for transport to a processing area for decontamination, preferably by autoclaving. |  |  |  |  |
| Sharps | Syringes which re-sheathe the needle, needleless systems, and other safety devices are used when appropriate. |  |  |  |  |
| Sharps | Broken glassware must not be handled directly by hand, but must be removed by mechanical means such as a brush and dustpan, tongs, or forceps. Containers of contaminated needles, sharp equipment, and brok en glass are decontaminated before disposal, according to any local, state, or federal regulations. |  |  |  |  |
| Sharps Injury Log | NA | NA | NA | The employer shall establish and maintain a sharps injury log for the recording of percutaneous injuries from contaminated sharps. The information in the sharps injury log shall be recorded and maintained in such manner as to protect the confidentiality of the injured employee. The sharps injury log shall contain, at a minimum: The type and brand of device involved in the incident, The department or work area where the exposure incident occurred, and an explanation of how the incident occurred. | NA |
| Spills | Spills and accidents that result in overt exposures to infectious materials are immediately reported to the laboratory director. Medical evaluation, surveillance, and treatment are provided as appropriate and written records are maintained. | Spills involving infectious materials must be contained, decontaminated and cleaned up by staff properly trained and equipped to work with infectious material. Spills involving infectious materials must be contained, decontaminated, and cleaned up by staff properly trained and equipped to work with infectious material. | Spills and accidents which result in overt exposures to organisms containing recombinant DNA molecules are immediately reported to the Institutional Biosafety Committee and NIH/ORDA. Reports to NIH/ORDA shall be sent to the Office of Recombinant DNA Activities, National Institutes of Health/ MSC 7010, 6000 Executive Boulevard, Suite 302, Bethesda, Maryland 20892-7010, (301) 496-9838. Medical evaluation, surveillance, and treatment are provided as appropriate and written records are maintained. | HIV and HBV Research Laboratories and Production Facilities: All spills shall be immediately contained and cleaned up by appropriate professional staff or others properly trained and equipped to work with potentially concentrated infectious materials. | A written procedure for the clean-up of all spills must be developed and followed. All spills, accidents and overt or potential exposures to infectious materials must be reported to the laboratory supervisor. A written record of such accidents and incidents should be maintained. |
| Training | The laboratory director ensures that laboratory and support personnel receive appropriate training on the potential hazards associated with the work involved, the necessary precautions to prevent exposures, and the exposure evaluation procedures. Personnel receive annual updates or additional training as necessary for procedural or policy changes. | The laboratory supervisor must ensure that laboratory personnel demonstrate proficiency in standard and special microbiological practices before working with BSL-2 agents. Spills involving infectious materials must be contained, decontaminated, and cleaned up by staff properly trained and equipped to work with infectious material. |  |  |  |
| Transport of contaminated materials | Cultures, tissues, specimens of body fluid s, or potentially infectious wastes are placed in a container with a cover that prevents leakage during collection, handling, processing, storage, transport, or shipping. | Materials to be removed from the facility for decontamination must be packed in accordance with applicable local, state, and federal regulations. Potentially infectious materials must be placed in a durable, leak proof container during collection, handling, processing, storage, or transport within a facility. | Contaminated materials that are to be decontaminated at a site away from the laboratory are placed in a durable leak-proof container which is closed before being removed from the laboratory. |  |  |
| C. BL2 Safety Equipment (Primary Barriers) |  |  |  |
| Biosafety Cabinets | Properly maintained biological safety cabinets, preferably Class II, or other appropriate personal protective equipment or physical containment devices are used. | All procedures involving the manipulation of infectious materials that may generate an aerosol should be conducted within a BSC or other physical containment devices | Biological safety cabinets (Class I or II) (see Appendix G-III-L, Footnotes and References of Appendix G) or other appropriate personal protective or physical containment devices are used. | Biological safety cabinets shall be certified when installed, whenever they are moved and at least annually. | Biological safety cabinets, to be used whenever there is an increased risk of airborne infection, and procedures with a high potential for producing aerosols are used. |
| Centrifuges | High concentrations or large volumes of infectious agents are used. Such materials may be centrifuged in the open laboratory if sealed rotor heads or centrifuge safety cups are used, and if these rotors or safety cu ps are opened only in a biological safety cabinet. | High concentrations or large volumes of infectious agents are used. Such materials may be centrifuged in the open laboratory using sealed rotor heads or centrifuge safety cups. | High concentrations or large volumes of organisms containing recombinant DNA molecules are used. Such materials may be centrifuged in the open laboratory if sealed beads or centrifuge safety cups are used and if they are opened only in a biological safety cabinet. |  |  |
| Coats, Gowns, or Uniforms. | It is recommended that laboratory coats, gowns, or uniforms be worn to prevent contamination or soiling of street clothes. Protective laboratory coats, gowns, smock s, or uniforms designated for lab use are worn while in the laboratory. This protective clothing is removed and left in the laboratory before leaving for non-laboratory areas (e.g., cafeteria, library, administrative offices). All protective clothing is either disposed of in the laboratory or laundered by the institution; it should never be taken home by personnel. | Protective laboratory coats, gowns, smocks, or uniforms designated for laboratory use must be worn while working with hazardous materials. Remove protective clothing before leaving for non-laboratory areas (e.g., cafeteria, library, administrative offices). Dispose of protective clothing appropriately, or deposit it for laundering by the institution. It is recommended that laboratory clothing not be taken home. | Laboratory coats, gowns, smocks, or uniforms are worn while in the laboratory. Before exiting the laboratory for non-laboratory areas (e.g., cafeteria, library, administrative offices), this protective clothing is removed and left in the laboratory or covered with a clean coat not used in the laboratory. | All personal protective equipment shall be removed prior to leaving the work area. Laboratory coats, gowns, smocks, uniforms, or other appropriate protective clothing shall be used in the work area and animal rooms. Protective clothing shall not be worn outside of the work area and shall be decontaminated before being laundered. When there is occupational exposure, the employer shall provide, at no cost to the employee, appropriate personal protective equipment such as, but not limited to, gloves, gowns, laboratory coats, face shields or masks and eye protection, and mouthpieces, resuscitation bags, pocket masks, or other ventilation devices. Personal protective equipment will be considered "appropriate" only if it does not permit blood or other potentially infectious materials to pass through to or reach the employee's work clothes, street clothes, undergarments, skin, eyes, mouth, or other mucous membranes under normal conditions of use and for the duration of time which the protective equipment will be used. | Laboratory coveralls, gowns or uniforms must be worn at all times for work in the lab. It is prohibited to wear protective laboratory clothing outside the laboratory, e.g. in canteens, coffee rooms, offices, libraries, staff rooms and toilets. Protective lab clothing that has been used in the lab must not be stored in the same lockers or cupboards as street clothing. |
| Equipment & Preventative Maintenance | None | Laboratory equipment should be routinely decontaminated, as well as, after spills, splashes, or other potential contamination. | None | Engineering and work practice controls shall be used to eliminate or minimize employee exposure. Where occupational exposure remains after institution of these controls, personal protective equipment shall also be used. Engineering controls shall be examined and maintained or replaced on a regular schedule to ensure their effectiveness. Equipment which may become contaminated with blood or other potentially infectious materials shall be examined prior to servicing or shipping and shall be decontaminated as necessary, unless the employer can demonstrate that decontamination of such equipment or portions of such equipment is not feasible. | Equipment such as autoclaves and biological safety cabinets must be validated with appropriate methods before being taken into use. Recertification should take place at regular intervals, according to the manufacturer’s instructions |
| Eye wash | An eyewash station is readily available. | An eyewash station must be readily available. |  | HIV and HBV Research Laboratories and Production Facilities: Each laboratory shall contain a facility for hand washing and an eye wash facility which is readily available within the work area. |  |
| Eye Wear | Face protection (goggles, mask, face shield or other splatter guard) is use d for anticipate d splashes or sprays of infectious or other hazardous materials to the face when the microorganisms must be manipulated outside the BSC. Protective eyewear should be worn for conduct of procedures in which splashes of microorganisms or other hazardous materials is anticipated. | Eye and face protection (goggles, mask, face shield or other splatter guard) is used for anticipated splashes or sprays of infectious or other hazardous materials when the microorganisms must be handled outside the BSC or containment device. Eye, face and respiratory protection should be used in rooms containing infected animals as determined by the risk assessment. |  | Masks, Eye Protection, and Face Shields. Masks in combination with eye protection devices, such as goggles or glasses with solid side shields, or chin-length face shields, shall be worn whenever splashes, spray, spatter, or droplets of blood or other potentially infectious materials may be generated and eye, nose, or mouth contamination can be reasonably anticipated. | Safety glasses, face shields (visors) or other protective devices must be worn when it is necessary to protect the eyes and face from splashes, impacting objects and sources of artificial ultraviolet radiation. |
| Gloves | Gloves should be worn if the skin on the hands is broken or if a rash is present. Alternatives to powdered latex gloves should be available. Gloves are worn when hands may contact potentially infectious materials, contaminated surfaces or equipment. Wearing two pairs of gloves may be appropriate. Gloves are disposed of when overtly contaminated, and removed when work with infectious materials is completed or when the integrity of the glove is compromised. Disposable gloves are not washed, reused, or used for touching “clean” surfaces (keyboard s, telephones, etc.), and they should not be worn outside the lab. Alternatives to powdered latex gloves should be available. Hands are washed following removal of gloves. | Gloves must be worn to protect hands from exposure to hazardous materials. Glove selection should be based on an appropriate risk assessment. Alternatives to latex gloves should be available. Gloves must not be worn outside the laboratory. In addition, BSL-2 laboratory workers should: a. Change gloves when contaminated, integrity has been compromised, or when otherwise necessary. Wear two pairs of gloves when appropriate. b. Remove gloves and wash hands when work with hazardous materials has been completed and before leaving the laboratory. c. Do not wash or reuse disposable gloves. Dispose of used gloves with other contaminated laboratory waste. Hand washing protocols must be rigorously followed. |  | HIV and HBV Research Laboratories and Production Facilities: Special care shall be taken to avoid skin contact with other potentially infectious materials. Gloves shall be worn when handling infected animals and when making hand contact with other potentially infectious materials is unavoidable. | Appropriate gloves must be worn for all procedures that may involve direct or accidental contact with blood, body fluids and other potentially infectious material or infected animals. |
| Warning Labels | None | None | None | A readily observable label in accordance with paragraph (g)(1)(i)(H) shall be attached to the equipment stating which portions remain contaminated. Warning labels shall be affixed to containers of regulated waste, refrigerators and freezers containing blood or other potentially infectious material; and other containers used to store, transport or ship blood or other potentially infectious materials These labels shall be fluorescent orange or orange-red or predominantly so, with lettering and symbols in a contrasting color. Labels shall be affixed as close as feasible to the container by string, wire, adhesive, or other method that prevents their loss or unintentional removal. |  |
| D. BL2 Facility Design (Secondary Barriers) |  |  |  |
| Autoclave |  |  | An autoclave for decontaminating laboratory wastes is available. | HIV and HBV Research Laboratories and Production Facilities: An autoclave for decontamination of regulated waste shall be available within or as near as possible to the work area. | At Biosafety Level 2, an autoclave or other means of decontamination should be available in appropriate proximity to the laboratory. |
| Backflow | None | None | None |  | A dependable supply of good quality water is essential. There should be no cross-connections between sources of laboratory and drinking-water supplies. An anti-backflow device should be fitted to protect the public water system. |
| Bench Tops | Bench tops are impervious to water and are resistant to moderate heat and the organic solvents, acids, alkalis, and chemicals used to decontaminate the work surface and equipment. | Bench tops must be impervious to water and resistant to heat, organic solvents, acids, alkalis, and other chemicals. | Bench tops are impervious to water and resistant to acids, alkalis, organic solvents, and moderate heat. |  | Bench tops should be impervious to water and resistant to disinfectants, acids, alkalis, organic solvents & moderate heat. |
| Biological Safety Cabinet | Install biological safety cabinets in such a manner that fluctuations of the room supply and exhaust air do not cause the biological safety cabinets to operate outside their parameters for containment. Locate biological safety cabinets away from doors, from windows that can be opened, from heavily traveled laboratory areas, and from other potentially disruptive equipment so as to maintain the biological safety cabinets’ air flow parameters for containment. | BSCs must be installed so that fluctuations of the room air supply and exhaust do not interfere with proper operations. BSCs should be located away from doors, windows that can be opened, heavily traveled laboratory areas, and other possible airflow disruptions. HEPA filtered exhaust air from a Class II BSC can be safely re-circulated back into the laboratory environment if the cabinet is tested and certified at least annually and operated according to manufacturer’s recommendations. BSCs can also be connected to the laboratory exhaust system by either a thimble (canopy) connection or a direct (hard) connection. Provisions to assure proper safety cabinet performance and air system operation must be verified. |  | HIV and HBV Research Laboratories and Production Facilities:Biological safety cabinets shall be certified when installed, whenever they are moved and at least annually. |  |
| Carpets and Rugs | Carpets and rugs in laboratories are not appropriate. | Carpets and rugs in laboratories are not permitted. |  |  |  |
| Cleanliness | The laboratory is designed so that it can be easily cleaned. | The laboratory should be designed so that it can be easily cleaned and decontaminated. | The laboratory is designed so that it can be easily cleaned. | Employers shall ensure that the worksite is maintained in a clean and sanitary condition. The employer shall determine and implement an appropriate written schedule for cleaning and method of decontamination based upon the location within the facility, type of surface to be cleaned, type of soil present, and tasks or procedures being performed in the area. | The laboratory should be kept neat, clean and free of materials that are not pertinent to the work. |
| Chairs | Chairs and other furniture used in laboratory work should be covered with a non-fabric material that can be easily decontaminated. | Chairs used in laboratory work must be covered with a non-porous material that can be easily cleaned and decontaminated with appropriate disinfectant. |  |  |  |
| Decontamination |  | A method for decontaminating all laboratory wastes should be available in the facility (e.g., autoclave, chemical disinfection, incineration, or other validated decontamination methods). |  |  |  |
| Doors | Laboratories should have doors for access control. Provide lockable doors for facilities that house restricted agents (as defined in 42 CFR 72.6). | Laboratory doors should be self-closing and have locks in accordance with the institutional policies. |  | HIV and HBV Research Laboratories and Production Facilities: Access doors to the work area or containment module shall be self-closing. | Doors should have vision panels, appropriate fire ratings, and preferably be self-closing. |
| Emergency Back-Up Generators | None | None | None |  | There should be a reliable and adequate electricity supply and emergency lighting to permit safe exit. A stand-by generator is desirable for the support of essential equipment, such as incubators, biological safety cabinets, freezers, etc., and for the ventilation of animal cages. |
| Eyewash | An eyewash station is readily available | An eyewash station must be readily available. |  |  |  |
| First Aid Rooms |  |  |  |  | First-aid areas or rooms suitably equipped and readily accessible should be available |
| Laboratory Furniture | Laboratory furniture is capable of supporting anticipated loading and uses. | Laboratory furniture must be capable of supporting anticipated loads and uses. | Laboratory furniture is sturdy. |  | Laboratory furniture should be sturdy. |
| Lighting | Illumination is adequate for all activities, avoiding reflections and glare that could impede vision. |  |  |  | Illumination should be adequate for all activities. Undesirable reflections and glare should be avoided. |
| Security | None | None | None |  | Laboratories and animal houses are occasionally the targets of vandals. Physical and fire security must be considered. Strong doors, screened windows and restricted issue of keys are compulsory. Other measures should be considered and applied,as appropriate, to augment Security Safety systems should cover fire, electrical emergencies, emergency shower and eyewash facilities. |
| Spaces | Spaces between benches, cabinets, and equipment are accessible for cleaning. Consider locating new laboratories away from public areas. | Spaces between benches, cabinets, and equipment should be accessible for cleaning | Spaces between benches, cabinets, and equipment are accessible for cleaning. |  | Facilities for eating and drinking and for rest should be provided outside the lab working areas. Ample space must be provided for the safe conduct of laboratory work and for cleaning and maintenance. Open spaces between and under benches, cabinets and equipment should be accessible for cleaning. |
| Sink | Each laboratory contains a sink for handwashing. | Laboratories must have a sink for hand washing. The sink may be manually, hands-free, or automatically operated. It should be located near the exit door. | Each laboratory contains a sink for hand washing. | Employers shall provide hand washing facilities which are readily accessible to employees HIV and HBV Research Laboratories and Production Facilities: Each work area shall contain a sink for washing hands and a readily available eye wash facility. The sink shall be foot, elbow, or automatically operated and shall be located near the exit door of the work area. | Hand-washing basins, with running water if possible, should be provided in each laboratory room, preferably near the exit door. |
| Storage | None | None | None |  | Storage space must be adequate to hold supplies for immediate use and thus prevent clutter on bench tops and in aisles. Additional long-term storage space, conveniently located outside the laboratory working areas, should also be provided. |
| Vacuum Lines | None | Vacuum lines should be protected with High Efficiency Particilate Air filters, or their equivalent. Filters must be replaced as needed. Liquid disinfectant traps may be required. | None | HIV and HBV Research Laboratories and Production Facilities: Vacuum lines shall be protected with liquid disinfectant traps and high-efficiency particulate air (HEPA) filters or filters of equivalent or superior efficiency and which are checked routinely and maintained or replaced as necessary. |  |
| Ventilation | There are no specific ventilation requirements. However, planning of new facilities should consider mechanical ventilation systems that provide an inward flow of air without recirculation to s paces outside of the laboratory. If the laboratory has windows that open to the exterior, they are fitted with fly screens. | There are no specific requirements on ventilation systems. However, planning of new facilities should consider mechanical ventilation systems that provide an inward flow of air without recirculation to spaces outside of the laboratory. | None | HIV and HBV Research Laboratories and Production Facilities: A ducted exhaust-air ventilation system shall be provided. This system shall create directional airflow that draws air into the work area through the entry area. The exhaust air shall not be recirculated to any other area of the building, shall be discharged to the outside, and shall be dispersed away from occupied areas and air intakes. The proper direction of the airflow shall be verified (i.e., into the work area). | In the planning of new facilities, consideration should be given to the provision of mechanical ventilation systems that provide an inward flow of air without recirculation. |
| Windows | If the laboratory has windows that open to the exterior, they are fitted with fly screens. | Laboratory windows that open to the exterior are not recommended. However, if a laboratory does have windows that open to the exterior, they must be fitted with screens. | If the laboratory has windows that open, they are fitted with fly screens. |  | If there is no mechanical ventilation, windows should be able to be opened and should be fitted with arthropod-proof screens. |
| Work Surfaces | Laboratory equipment and work surfaces should be decontaminated with an effective disinfectant on a routine basis, after work with infectious materials is finished, and especially after overt spills, splashes, or other contamination by infectious materials. Contaminated equipment must be decontaminated according to any local, state, or federal regulations before it is sent for repair or maintenance or pack aged for transport in accordance with applicable local, state, or federal regulations, before removal from the facility. | None | None | HIV and HBV Research Laboratories and Production Facilities: The surfaces of doors, walls, floors and ceilings in the work area shall be water resistant so that they can be easily cleaned. Penetrations in these surfaces shall be sealed or capable of being sealed to facilitate decontamination. | Work surfaces must be decontaminated after any spill of potentially dangerous material and at the end of the working day. |
| Walls | None | None | None |  | Walls, ceilings and floors should be smooth, easy to clean, impermeable to liquids and resistant to the chemicals and disinfectants normally used in the laboratory. Floors should be slip-resistant. |