# ON THE MECHANCAL BEHAVIOR OF NITINOL UNDER MULTIAXIAL LOADING CONDITIONS AND IN SITU SYNCHROTRON X-RAY DIFFRACTION

# V. Imbeni*,**, A. Mehta***, S.W. Robertson*, T.W. Duerig****, A. Pelton**** and R.O. Ritchie*

* Materials Sciences Division, Lawrence Berkeley National Laboratory, and Department of Materials Science and Engineering, University of California, Berkeley

** SRI International and The International Organization on Shape Memory and Superelastic Technologies

*** Stanford Synchrothron Radiation Laboratory, University of Stanford, Palo Alto

**** Nitinol Devices and Components, Fremont, California

# ABSTRACT

Although the use of Nitinol in the biomedical industry has been very successful, there have been distinct problems due to the difficulty of being able to accurately compute stresses, particularly in the presence of complex in vivo loading patterns. Constitutive laws for shape-memory alloys subjected to multiaxial loading, which are based on direct experimental observations, are generally not available in the literature. Accordingly, in the present work, tension–torsion tests are conducted on Nitinol thin-walled tubes (thickness/radius ratio of 1:10) using various loading/unloading paths under isothermal conditions. The experimental results showed significant variations in the mechanical response along the two loading axes. To fully understand the macroscopic and microscopic aspects of the transformation in situ multiaxial (tension-torsion) loading experiments were also performed at the SSRL (Stanford Synchrothron Radiation Laboratory) in order to monitor the material’s state of transformation and identifying the nature, volume fraction of phases present and the elastic strain in them as a function of stress and stress state.

Keywords: Nitinol, multiaxial, torsion, X-ray, synchrotron, Lüders bands, R-phase, in situ

# INTRODUCTION

Most of the mechanical testing on polycrystalline Nitinol found in the literature has been performed on wires and is thus one-dimensional. As a result, most phenomenological constitutive models are based on uniaxial data, oftentimes extended to three-dimensions in an ad hoc fashion. Although there are several one-dimensional models, the lack of such constitutive relationships for these materials presents a real limitation in the design and in vivo reliability of the many medical devices manufactured from Nitinol. Design challenges have arisen from the difficulty of measuring and computing stresses and strains in the presence of complex in vivo loading during service life, as well as the incomplete understanding of Nitinol’s constitutive behavior (i.e., the relationship between stress and strain) under these conditions.

Very little information exists on the multiaxial loading/unloading of Nitinol and even less on tubes; to our knowledge, only three previous studies, all on thick-walled tubing, are available. [1,2,3] The present work focuses on the biaxial testing of thin-walled tubes, the starting material for the manufacture of endovascular stents and chosen to minimize the gradient in the torsional strain along the radial direction. Three distinct multiaxial stress-loading paths are explored under isothermal conditions.

The resulting data are also reproduced in an in situ testing rig at the Stanford Synchrotron Radiation Laboratory (SSRL), while simultaneously using synchrotron x-ray diffraction to discern the differences in phases present between tension and torsion observed during multiaxial loading. The novel in situ synchrotron x-ray experiments provided insight on the martensitic phase transformation associated with superelastic transformations under complex strain states. The long-term goal of this work is to incorporate the ‘macroscopic’ deformation data measured under tension/torsion loading with the ‘microscopic’ in situ x-ray diffraction tension/torsion results in order to develop a three-dimensional constitutive model for superelasticity in Nitinol.

# EXPERIMENTAL PROCEDURES

# MATERIALS

Nitinol tubing (Ti 49.2 at.%, Ni 50.8 at.%) was received from NDC (Fremont, CA), with a 4.64 mm outer diameter and 0.37 mm thickness. The tubing was cut into 75 mm long specimens and ground down along the 25 mm long center test section to an outer diameter of 4.3 mm in order to obtain a waisted shape and minimize end effects during testing.  The wall thickness in the gauge section was reduced to 0.2 mm, resulting in a thickness-to-radius ratio of approximately 1:10. The specimens were subsequently heat-treated at 485°C in air for 5 min and then ice water-quenched to bring the austenite finish temperature (Af) slightly below room temperature. This treatment produces a complex microstructure consisting of equiaxed low-angle subgrains, ~1 m in size, within irregularly shaped parent grains, ~25 m in size.

The transformation temperatures of the heat-treated samples were characterized by differential scanning calorimetry (DSC) and were found to be As = -6.4(C, Af  = 18.1(C, and Ms = -51.6(C, Mf = -87.4(C, respectively. X-ray diffractometry conducted at SSRL was used to observe the phases present before and after the testing.

# MULTIAXIAL MECHANICAL TESTING

As the shape change during the austenite-to-martensite transformation can create slippage problems in the grips, reliable tests could not be performed with straight tubes. Therefore tubes with a gauge section of reduced diameter were used in order to minimize end effects at the grips and to lower the stress in the gripped section. Also, a split-grip design was employed; a hardened steel pin was inserted into the tube throughout the grip region which allowed application of sufficient lateral gripping force to inhibit slippage during testing. Strain rates in the range 10-5–10-4 s-1 were used for both the tensile and torsional loading in order to avoid strain rate effects. The loading rates used in the experiments were also shown to be sufficiently low to suppress significant temperature variations during testing [4] such that all experiments took place under essentially isothermal conditions. More details on the mechanical test setup can be found in ref. 6.

## Loading Programs

In order to investigate thoroughly the complex nature of the forward and reverse transformation process and its effect on the constitutive properties of polycrystalline Nitinol, three specific loading/unloading programs (referred to as Types I–III) were investigated, as outlined below:

Type I:  (Load in tension followed by torsion, unload in reverse order)

1. Tensile loading to a particular strain level; hold during segments 2 and 3; 2. Torsional loading to 2% shear strain; 3. Unload shear axis to zero strain; 4. Unload tensile axis to zero strain.

Type II:  (Load in torsion followed by tension, unload in reverse order)

1. Torsional loading to 2% shear strain; hold during segments 2 and 3; 2. Tensile loading to a particular strain level; 3. Unload tensile axis to zero strain; 4. Unload shear axis to zero strain.

Type III:  (Load and unload in simultaneous tension–torsion)

1. Loaded simultaneously in tensile and torsional axes.  Rates for each axis were independently controlled so that the end strain was reached simultaneously; 2. Unloaded tensile and torsional axes simultaneously to zero strain.

A series of specific loading paths were considered for Types I–III, involving tensile strains ranging from 0% to 6.0% and the maximum attainable shear strain limited to 2% (due to buckling of the thin-walled structure). As the forward transformation is initiated at ~1% shear strain, this limit did not substantially restrict the mapping of the loading in the shear strain dimension.  A minimum of two repetitions were carried out for each loading path considered.

## IN SITU X-RAY SYNCHROTRON DIFFRACTION

For the in situ investigation of phases and the elastic strain of Nitinol tubes, the high-resolution x-ray, synchrotron based instrument, such as the one that was used at SSRL (BL2-1), offers many advantages over a conventional parafocusing x-ray instrument based on a sealed-tube as the source of radiation.  Such an instrument utilizes the high brightness of the x-ray beam and a perfect crystal analyzer (float zone Si 111) to easily achieve angular resolution sufficiently high to detect changes in lattice parameters smaller than 10 picometers.  Not only was the ultra-high resolution very useful in detecting the R-phase in Nitinol (Figure 1), but it also enabled us to very accurately measure the elastic strain at a local level in our in situ loading tests.  The high brightness of the synchrotron beams provided two addition advantages for our measurements.  These bright x-ray beams can be focused to a spatial spot size of 0.5x2 mm (FWHM), which was further reduced by using slits.  This enabled us to probe the phenomenon of inhomogeneous transformation of the martensite during loading of the tubes in tension.  Furthermore, the high degree of collimation of the synchrotron beam facilitates using a parallel-beam diffraction geometry. In a parallel-beam geometry the location of the Bragg peak is not dependent on the deviation of the diffracting surface of the sample from the center of the diffractometer, and therefore it is ideal for investigating a curved sample (a thin-walled tube) in an in situ deformation rig.

Tension and torsion tests were performed in situ on the same samples employed for the multiaxial loading described above. Due to the extreme complexity of the non-proportional loading tests, where tension and torsion were combined, only the pure tension and torsion experiments where reproduced in situ using a custom designed rig.  The tubes were placed in the in situ rig and gradually loaded up to 6% strain in tension and 2.5% in torsion (until buckling occurred).  X-ray diffraction patterns were collected at fixed strains along the gauge length of the tube, and patterns were collected at 1 mm intervals along the tube. This procedure was chosen in order to study the nucleation and growth of martensite spiral bands (similar in appearance to Lüders bands), which is observed during loading of tubes in tension (Figure 2).

Figure 1: Nitinol diffraction patterns: the characteristic patterns of austenite, R-phase and martensite are shown. This pattern was collected on a fully annealed sample to highlight the effect of the R-phase.

Figure 2: Spiral martensite bands appear during tensile loading of an electropolished tube, similar in appearance to Lüders bands observed in mild steel

# RESULTS AND DISCUSSION

# MULTIAXIAL MECHANICAL TESTING

Results of the experiments conducted in this study are presented in Figs. 3-8 as plots of equivalent true (Cauchy) stress eq versus equivalent referential (Lagrangian) strain eq, i.e., eq  = ((t2 + 3 s2) and eq  = ((t2 + 4/3s2), where t and t (respectively, s and s) denote the tensile (respectively, shear) stress and strain. Within acceptable experimental error margins, the unloading of all stress–strain curves passed back through the origin, indicating that the transformation was reversible and displays the superelastic effect independently of the loading path.

# Uniaxial tension and torsion

Pure tension results are shown in Figure 3 and reveal behavior typical of shape-memory alloys in the superelastic regime. The repeatability of the tests appears to be adequate with the onset of the transformation varying between 1.1% and 1.2% strain and the transformation plateau stresses lying between 370 and 380 MPa. Under pure torsional loading (Figure 4), the transition from the elastic loading region to the transformation plateau region is less sharp.  There is a lack of closure in the loading/unloading curve due to the slippage of the axial extensometer during torsion; this is deduced by simple comparison with the corresponding shear (as opposed to equivalent) stress–strain diagram where the closure is accurately reproduced.

Figure 3:  Equivalent stress–strain plot showing the repeatability of three separate uniaxial tension tests loaded up to 6% strain.

Figure 4:  Equivalent stress-strain plot showing pure torsion loaded up to 2% shear

Tension followed by torsion (Type I)

Results of Type I loading experiments involving tension up to a maximum of 6% strain, followed by torsional loading of 2% strain (tension is held constant) and reverse-order unloading are presented in Figure 6. Even with this relatively simple loading path, the material response is quite complex. The stress-induced austenite-to-martensite transformation is effected by the formation of martensitic structures which correspond to system energy minimizers known as variants.  Although 24 variants of the less symmetric martensitic phase can be formed by the same crystal of parent austenite, it has been experimentally observed that only a few variants are typically active and fully resolve the stress/shape change.[1,5]At low tensile strains (<1%), the overall behavior is similar to that of pure torsion; at higher tensile strains, a sharp increase is observed in the stress upon superposition of torsion, which is indicative of the generation of a new set of martensitic variants which are energetically preferable for torsional (rather than the tensile) deformation. In addition, the hysteresis in the torsional loading/unloading loop appears to decrease progressively with increasing applied tensile strain, which is consistent with the decreasing volume fraction of variants associated with torsion. At all levels of applied tensile strain, the end of the torsional unloading segment is marked by an increasing equivalent stress with decreasing applied equivalent strain. The micromechanical origin of this mechanical response is not presently understood.  However, it does qualitatively indicate that the material has retained a microstructural ‘‘memory’’ of its tensile transformation state.

Figure 5:  Equivalent stress–strain plots in tension (0%, 0.7%, 1.10%, 1.5%, 2%, 3%, 6%), followed by torsion (2%) and reverse unloading (Type I).

Torsion followed by tension (Type II)

In this case, torsion is applied corresponding to 2% strain, followed by tension up to 5.8% strain. Unloading is subsequently effected in the reverse order. Type II loading paths differ substantially from Type I paths, as seen in Figure 6: here the transformation plateau that follows the torsional loading occurs at monotonically decreasing stress levels. In addition, upon unloading, the ‘‘elastic’’ region is followed by reverse transformation along a plateau at monotonically increasing stress levels up to the point that corresponds to the end of the tensile unloading.  Again, the material appears to exhibit ‘microstructural’ memory of the transition from proportional to non-proportional loading.

Figure 6:  Equivalent stress–strain plots in torsion (2%), followed by tension (0%, 0.7%, 1.05%, 3%, 5.8%) and reverse unloading (Type II).

Simultaneous tension–torsion (Type III)

The results of simultaneous proportional tension–torsion at up to 6.0% tensile strains and at 2.0% shear are shown in Figure 7. All loading paths produce curves that are qualitatively similar to those obtained under pure tension or torsion. In particular, smooth loading and unloading transformation plateaus are obtained at approximately constant stresses. This observation appears to support the conjecture that sharp changes in the stresses during non-proportional loading are effected by the generation of new sets of martensitic variants, while existing variants lead to higher stresses due to their energetic incompatibility with the current loading state.

Figure 7:  Equivalent stress–strain plots in simultaneous tension (0%, 0.7%, 1.5%, 3%, 6%) and torsion (2%) with simultaneous unloading (Type III).

Figure 8:  Comparison of equivalent stress and strain behavior between loading paths at peak strains of 3.0% in tension and 2.0% in torsion.

## Comparison of results from various loading path types

Figure 8 compares Types I–III loading paths for 3.0% tensile strains and shearing strain of 2%. At low tensile strain levels, the loading paths are quite similar. At moderate tensile strains, Type I and Type II loading generates widely differing “non-smooth” stress–strain curves, owing to the sharp transitions in variant formation at the transformation plateau. At a high tensile strain level, the Types II and III paths yield remarkably similar curves, which can be explained by the predominantly proportional nature of both paths during the transformation process. More detail can be found in ref. 6.

## Analysis of results: qualitative features

The experiments demonstrate that the development of a load-dependent microstructure could be responsible for the non-smooth stress–strain curves under non-proportional loading of Types I and II. In addition, the alloy appears to exhibit memory of its forward transformation history throughout the loading process, as illustrated by the reverse transformation paths followed in the non-proportional loading cases. Another important feature observed in the experiments is that the forward transformations in tension and torsion follow markedly different paths. In particular, and contrary to the case of initial yielding in metals, the transformation stress in torsion is consistently higher than in tension.

In order to investigate further the different deformation mechanisms in tension and torsion, electropolished specimens were loaded according to the Types I and II protocols. During Type I loading, a spiral martensite band similar in appearance to a Lüders band [7] nucleated and grew. Its formation was observed as soon as the tensile stress reached the plateau value of ~400 MPa at 1% tensile strain, as shown in Figure 3; the spiral band remained visible during the torsion cycle, where the tensile strain is held constant.  The martensite band only disappeared at the end of the tension plateau where the sample was fully martensitic as confirmed by the in situ x-ray experiments. In contrast, no Lüders bands were observed during Type II loading (even during the tension cycle applied at a fixed shear stress). The absence of Lüders band in torsion was also verified by the in situ x-ray diffraction experiments.  It is believed that Lüders band formation during the tensile cycle facilitates the deformation, consistent with the higher transformation stresses required in torsion where such bands do not appear to form (at least up to 2% shear).  This is also supported by the simultaneous loading tests (Type III), where the plateau stress decreases with increasing tensile strain (Figure 7).  Another interesting characteristic of the stress–strain curves is that in tension (Figure 3), there is a significant reduction in the elastic modulus at very low strains (~0.3%), which may be attributed to the presence of the R-phase; interestingly, the torsion experiments show no such reduction in modulus (Figure 4), as again verified in the in situ torsion experiments.

## Constitutive modeling

The tension–torsion experiments described above were simulated by a finite element analysis using a three-dimensional superelasticity model in finite deformations. The model predicts phase transformation potentially involving all 24 habit plane martensitic variants [8.9]. The computational method for solving the constitutive equations associated with the above model is discussed in a companion paper (Jung et al., SMST 2003 proceedings). The model predicts reasonably well the stress response in both tension and torsion, due to the inclusion of different martensitic variants that can nucleate under non-proportional, non-simultaneous tensile and torsional loading. This feature, which is clearly exhibited in the experiments and replicated in the numerical simulation, has been difficult to capture using previous phenomenological constitutive models [9].

## MICROSCOPIC IN SITU X-RAY SYNCHROTRON DIFFRACTION

Through in situ multiaxial loading of Nitinol in tension and torsion, quantitative estimates of the nature and volume fraction of phases was assessed as a function of stress/strain state and of position along the tube. Also, the perpendicular elastic strain field (Poisson strain) was measured from the shifts of the characteristic peaks upon loading as a function of position along the tube, stress and strain. The perpendicular strain varies with applied strain and with phase (austenite/martensite). The volume fraction of martensite present is calculated from the area under the peaks.

### Uniaxial tension and torsion

In uniaxial tensile loading, the stress-induced martensitic transformation is inhomogeneous and occurs by initiation and growth of a spiral martensite band (Lüders bands) with a sharp austenite-martensite interface (Figures 9,10).  The Lüders bands observed during the tensile loading cycle were fully characterized in situ. Their presence was confirmed by x-ray data; outside the Lüders bands the material is fully austenitic whereas within the bands it is fully martensitic (Fig 9,10). The bands are actually a spiral that wraps around the tube.

Fig 9:  Poisson elastic strain vs. position along the tube at various equivalent tensile strains

Fig 10: Volume fraction of martensite vs. position along the tube at various equivalent tensile strains

Fig 11:  Poisson strain vs. position along the tube in torsion

Fig 12: Volume fraction of martensite vs. position along tube in torsion

Fig 13: Volume fraction of martensite vs. equivalent strain in tension and torsion

Fig 14:  Volume fraction of martensite vs. Poisson strain (perpendicular elastic strain).

The tubes behaved very differently under torsion. Both in situ x-ray diffraction and surface morphology analysis indicated that the deformation is homogeneous across the whole tube (Figures 11-12). In torsion, the volume fraction of martensite and the Poisson strain (as a function of equivalent strain) varied uniformly (Figures13-14), with no Lüders band formation. The experiments revealed a surprising path dependance of Nitinol loading that is bound to affect the in vivo behavior where multiaxial stresses are present.

# SUMMARY AND CONCLUDING REMARKS

Tension–torsion experiments on thin-walled superelastic Nitinol tubes have demonstrated a rich spectrum of mechanical behavior involving significantly different characteristics in each of the two loading axes.  Indeed, the behavior of Nitinol under multiaxial loading is extremely complex yet displays remarkable microstructural memory. This behavior provides important insights into the stress-induced phase transformation under multiaxial loading commonly encountered in a wide array of engineering applications, in particular in the biomedical industry. The experiments suggest that the characterization and modeling of the microstructure is of paramount importance in understanding the phenomenology of shape-memory alloys. This characterization and modeling must include both mechanics- and microstructural-based understanding.

Under tensile loading, the phase transformation is highly heterogeneous and involves Lüders band formation; corresponding behavior in torsion is essentially homogeneous without the formation of Lüders bands.  In addition, the transformation stress is 50-70 MPa higher in torsion compared to tensile loading.  We surmise that load-dependent microstructures (i.e., variants of martensite dictated by the loading conditions) may be responsible for the distinctly differing stress-strain behavior in tension and torsion.

# ACKNOWLEDGEMENTS

This work was supported by the National Science Foundation Grant no. CMS-9800006 with the University of California at Berkeley. Special thanks are due to NDC for additional financial support and for supplying the Nitinol tubing, to Dr. J. H. McNaney for his initial work on the tension/torsion testing, and to Prof. Panos Papadopoulos and Youngjean Jung for many helpful discussions. In situ x-ray diffraction measurements were carried out at the Stanford Synchrotron Radiation Laboratory (SSRL), a national user facility operated by Stanford University on behalf of the U.S. Department of Energy

# REFERENCES

- S. Miyazaki, K. Otsuka, Wayman, C.M, Acta Metall. Vol. 37, 1989, pp.1873-1890.

- D. Helm, , Haupt, P, In: C.S. Lynch, (Ed.), Smart Structures and Materials. Bellingham, SPIE, 4333, 2001, pp. 302–313.

- Lim, T.J., McDowell, D.L., ASME J. Eng. Mater. Tech. 121, 9–18.

- Lim, T.J., McDowell, D.L., J. Intell. Mater. Sys. Struct. Vol. 6, 1999, pp. 817-830

- Gall, K., Sehitoglu, H., Chumlyakov, Y.I., Kireeva, I.V, Acta Mater. Vol. 47 (4), 1999, pp. 1203-1217.

- McNaney, J.H., Imbeni V., Jung, Y., Papadopolous, P. and R.O. Ritchie, Mech. Mater., Vol. 35 (10), 2003, pp. 969-986.

- Qing-Ping Sun, Zhi-Qi Li, Int. J. of Solids and Structures Vol. 39, 2002, pp. 3797-3809

- Ling, H.C., Karlow, R, Metall. Trans. A 12A, 1981, pp. 2101-2111.

- Jung, Y., Papadopoulos, P., Ritchie, R.O., Int. J. Numer. Methods Engin. 2003, in press.